## Slide 1
Bioresorbale Stents
Just an Alternative or One-Way Street?
Georgios I. Papaioannou, MD, MPH, FACC, FSCAI
Director, Cardiac Catheterization Laboratory
Athens Medical Center
29/10/2015

## Slide 2: Bioresorbale Drug-Eluting Stents (BRS)
Breakthrough technology (after biodegradable polymers)
Stents disappear after their useful function (no local inflammatory reaction)
Restore vasomotor tone and endothelialization
Increase vessel diameter (positive remodeling)
Improve coronary physiology
Potentially decrease anginal burden

## Slide 3: Bioresorbale Drug-Eluting Stents (BRS)
Not currently all patients are candidates and lesions are suitable
Stent deployment requires optimal pre and post dilatation
Low threshold of intracoronary imaging techniques
Higher rate of early post-procedural stent thrombosis

## Slide 4: Optimal BRS
Optimal short and mid term radial support with thin struts
Adequate deliverability, handling and flexibility for insertion
Consistency of drug elution
Integrity during resorption

## Slide 5: Strategies of developing BRS
Backbone comprised of L-lactic acid polymer
  ABSORB stent (BVS): Everolimus-eluting stent (CE Mark, Abbot Vascular)
  DESolve stent: Myolimus/Novolimus-eluting stent (CE Mark, Elixir Medical Corporation)
  ART stent: No eluting drug (CE Mark, Terumo Corporation)
Magnesium-based scaffolds (with rare earth metals)
  AMS/DREAMS: Paclitaxel/Sirolimus-eluting stent (Biotronic)
Other (Desaminotyrosin, Polylactic anhydrate)
  REVA BRS/ReZolve: Paclitaxel/Sirolimus-eluting stent (Reva Medical)
  Ideal Biostent: Sirolimus-eluting stent (Xenogenics)

## Slide 6: Types of BRS
A: Igaki-Tamai stent; B: ABSORB stent; C: DESolve stent; D: DREAMS Magesium alloy; E: ReZove BRS; F: Ideal BioStent

## Slide 7: Igaki-Tamai Stent
PLLA-based BRS
Self-expandable when contrast dye heated at 800 C
Initial results promising and vessel diameter increased at 3-year follow-up
Requires 8F Catheter
Heated contrast dye may cause injury

## Slide 8: Absorb BVS
PLLA-based stent
150 μm strut thickness (c/w 90 in current DES)
1:1 mixture of poly-D,L-lactic acid and everolimus
Degrades over time to H2O and CO2
Full hydrolytic degadration up to 3 years

## Slide 13: ABSORB III – 2008 Patients
Outcome | Absorb (1322) (%) | Xience (686) (%) | p
Target lesion failure | 7.8 | 6.1 | 0.16
Cardiac death | 0.6 | 0.1 | 0.29
Target vessel MI | 6 | 4.6 | 0.18
Ischemia-driven target lesion revascularization | 3 | 2.5 | 0.5
Ischemia-driven target vessel revascularization | 5 | 3.7 | 0.18
Patient-reported angina | 18.3 | 18.4 | 0.93
Definite or probable stent thrombosis | 1.5 | 0.7 | 0.13

## Slide 14
Rates of Stent Thrombosis after implantation of BRS vs BMS

## Slide 15: DESolve BRS
PLLA based-stent
150 μm strut thickness
Novolimus eluting drug
Degrades to H2O and CO2 in 1 year
Device resorption time is 2 years
Wider range of expansion, reduced strut fracture, self-correction of minor malapposition

## Slide 16: DESolve Nx Clinical Trial Design
Primary Endpoint: 6-month in-scaffold late lumen loss
Secondary Endpoints:
  Clinical: Major Adverse Cardiac Events (cardiac death, target vessel MI, and clinically indicated TLR), Scaffold thrombosis
QCA: In-segment late lumen loss, binary restenosis, and percent diameter stenosis
  IVUS: In-scaffold percent volume obstruction, malapposition
  OCT: In-scaffold percent obstruction, strut coverage
  MSCT: Percent diameter stenosis, lumen area
13 International Sites
Europe, New Zealand and Brazil
126 patients
30d
6mo
5yr
4yr
3yr
2yr
Clinical MACE
Angiographic
IVUS, OCT, MSCT (subset)
1yr
Single de novo Coronary Artery Lesion
Reference vessel diameter: 2.75-3.0mm
Lesion length: <12mm, DAPT 12 months
3.0, 3.25, 3.5mm diameters; 14 and 18mm lengths
Notes: The DESolve NX study demonstrated consistent findings through multiple imaging modalities IVUS, OCT, QCA or multi scan CT.

## Slide 17: DESolve Nx showed excellent acute performance and low events through 2 years
EKE 2015

## Slide 18: Case Example 1 - Multi-modality Imaging
SA = 7.73mm2
SA = 9.50mm2
SA = 6.09mm2
SA = 6.48mm2
SA = 7.00mm2
12 mos FU
IVUS/OCT
6 mos FU
MSCT
Pre Procedure
Post-PCI
Post-PCI
6 mos FU
Mean Scaffold Diameter = 3.1mm
Angio’s
Notes: Throughout different imaging modalities and timeframes the results remain consistent. Reference: Dr Verheye DESolve First in Man

## Slide 19: ART BRS
PLLA or PDLA based BRS
No eluting drug – next generation with eluting drug (Terumo)
Programmed dismantling at 3 months and resorption in 2 years
CE Mark based on ARTDIVA study 6 months results

## Slide 20: DREAMS 2G
Magnesium-based BRS
120-150 μm strut thickness
Electronegative charge emerges during degradation with antithrombotic function
Paclitaxel to sirolimus (2G) eluting drug
Resorption time up to 12 months

## Slide 24: Safety and performance of the second-generation drug-eluting absorbable metal scaffold in patients with de-novo coronary artery lesions (BIOSOLVE-II): 6 month results of a prospective, multicentre, non-randomised, first-in-man trial
Between Oct 8, 2013, and May 22, 2015, we enrolled 123 patients with 123 coronary target lesions. At 6 months, mean in-segment late lumen loss was 0·27 mm (SD 0·37), and angiographically discernable vasomotion was documented in 20 (80%) of 25 patients. Intravascular ultrasound assessments showed a preservation of the scaffold area (mean 6·24 mm2 [SD 1·15] post-procedure vs6·21 mm2 [1·22] at 6 months) with a low mean neointimal area (0·08 mm2 [0·09]), and optical coherence tomography did not detect any intraluminal mass. Target lesion failure occurred in four (3%) patients: one (<1%) patient died from cardiac death, one (<1%) patient had periprocedural myocardial infarction, and two (2%) patients needed clinically driven target lesion revascularisation. No definite or probable scaffold thrombosis was observed.

## Slide 25: Conclusions
BRS represent an advancement in the interventional treatment of coronary disease
Current data are slightly inferior with respect to device success, recoil, MACE, lumen areas and TLR
Some possible benefits are only demonstrated in animal testing and small human cohorts
Data on various type of patients and lesions are limited
Technical considerations of deployment and imaging
Optimal duration of double antiplatelet therapy is unclear
However, current BRS limitations will likely be resolved in the near future